i1PRO 3	Kalibracja i profilowanie na wyższym poziomie

Spectral Engine

i1Pro 3: i1® technology (holographic diffraction grating with 128 pixel diode array) with built-in wavelengths check

| Spectral Range | 380 - 730 nm |
| --- | --- |
| Physical Sampling Interval | 3.5 nm |
| Optical Resolution | 10 nm |
| Spectral reporting | 380 nm to 730 nm in 10 nm steps |
| Sampling Rate (Hz) | 400 measurements per second |

Optics

| Measurement Geometry | 45º/0º ring illumination optics, ISO 13655:2017 |
| --- | --- |
| Measurement Aperture | 4.5 mm (0.18”) diameter |
|  | (effective measurement aperture during scanning is |
|  | depending on the patch size and measurement speed) |
| Light Source | Full-Spectrum LED Light Source |

Reflectance Measurement

| Measurement Condition | M0 - UV included - ISO 13655:2017 |
| --- | --- |
|  | M1 - D50 - ISO 13655:2017 |
|  | M2 - ISO 13655:2017 |
| OBC | Optical Brightener Compensation (OBC) with |
|  | i1Profiler software |
| Calibration | Manual on external ceramic white reference |
| Measurement Background | White, ISO 13655:2017; for measurements |
|  | on backup board |
| Minimal Patch Size In Spot Mode | 6 x 6 mm (0.24” x 0.24”) (Width x Height) |
| Minimal Patch Size In Scanning Mode | 6 x 7 mm (0.24” x 0.28”) (Width x Height) with i1iO |
|  | 7 x 7 mm (0.28” x 0.28”) (Width x Height) with sensor ruler |
|  | 10 x 7 mm (0.40” x 0.28”) (Width x Height) without sensor |
|  | ruler |
| Maximum Scan Length | 260mm (10.24”) |
| Inter-Instrument-Agreement | 0.3 dE00 avg, 0.8 dE00 max (deviation from X-Rite manu- |
|  | facturing standard at a temperature of 23ºC (73.4ºF) on 12 |
|  | BCRA tiles (D50, 2º)) |
| Short-term Repeatability | 0.05 dE00 on white (D50,2°, mean of 10 measurements |
|  | every 3 s on white) |

by

| i1PRO 3 | Taking Calibration and Profiling to the Next Level |
| --- | --- |
|  |  | Emission Measurement |  |  |
|  |  | Measurement Range | 0.2 - 5000 cd/m2 on a typical LCD-Monitor |
|  |  | Short-term Repeatability | x,y ±0.002 typ. (5000 K, 80 cd/m2) |
|  |  | Ambient Light Measurement |  |  |
|  |  | i1Pro 3 | Spectral irradiance [mW/nm/m2], illuminance [lux] |
|  |  |  | cosine corrected diffusor light measurement head |
|  |  | Interface |  |  |
|  |  | USB 1.1 |  |  |
|  |  | Power Supply |  |  |
|  |  | Device powered by USB. No additional |  |  |
|  |  | charger or battery required. USB 1.1 |  |  |
|  |  | high power device. |  |  |
|  |  | Operating Conditions |  |  |
|  |  | Temperature | 10°C (50°F) – 35°C (95°F) |
|  |  | Humidity | 0% – 80% non-condensing |
|  |  | Storage Conditions |  |  |
|  |  | Temperature | 0° to 45°C |
|  |  | Physical Dimensions |  |  |
|  |  | i1Pro 3 Device | Length 162 mm, width 69 mm, height 64 mm |
|  |  |  | (6.4” x 2.7” x 2.5”) |
|  |  | i1Pro 3 Ruler | 337 mm x 102 mm (13.3” x 4.0”) |
|  |  | i1 Pro 3 Backup Board | 355 mm x 265 mm (14.0” x 10.4”) folded or |
|  |  |  | 355 mm x 400 mm (14.0” x 15.7”) unfolded |
|  |  | Weight i1Pro 3 Device | 285 g (10.1 oz) |

For more information, please visit xrite.com or contact your local X-Rite sales representative.

by

X-Rite is either a registered trademark or trademark of X-Rite, Incorporated in the United States and/or other countries. PANTONE©, PantoneLIVE and other Pantone trademarks are the property of Pantone LLC. All other trademarks or registered trademarks are the property of their respective owners. © X-Rite, Inc. 2019. All rights reserved.

L7-718 (12/19)